## Slide 1: Towards accurate wearable data for clinical decision making
Dr Peter H. Charlton
University of Cambridge
City, University of London
Photo by Luke Chesser on Unsplash

## Slide 2
7,000
Strokes prevented
Lives saved p.a.
2,000
425,000
Additional diagnoses
If atrial fibrillation was adequately treated in England:
Stroke Association, “State of the Nation,” 2017.
Public Health England, “Atrial fibrillation prevalence estimates in England …”, 2015.
Pexels, https://pixabay.com/photos/man-heartache-chest-pain-hurt-pain-1846050/ (Pixabay Licence)
Notes: Source: https://www.stroke.org.uk/professionals/atrial-fibrillation-information-and-resources ; “It's estimated that if AF was adequately treated, around 7,000 strokes would be prevented and over 2,000 lives saved every year in England alone.”

## Slide 3
Photo by Luke Chesser on Unsplash

## Slide 4
Source: Zenicor Press Material, https://zenicor.com/press-material/

## Slide 5
1. Introduction to wearables
2. Our contributions
Photo by Luke Chesser on Unsplash
3. Perspectives
Photo by Dario Valenzuela on Unsplash

## Slide 6
1. Introduction to wearables
2. Our contributions
Photo by Luke Chesser on Unsplash
3. Perspectives
Photo by Dario Valenzuela on Unsplash

## Slide 7
A Fitness tracker which acquires photoplethysmography (PPG) and accelerometry (Accel.) signals
A Smartwatch which acquires electrocardiography (ECG) and accelerometry (Accel.) signals
Pulse waves used to:
- estimate heart rate
- identify an irregular pulse
Steps used to:
- estimate step count
Heart beats. Signal used to:
- estimate heart rate
- (potentially) diagnose
arrhythmias
Peter H Charlton, https://doi.org/10.5281/zenodo.798234 (CC BY 4.0)
Source: NARA & DVIDS Public Domain Archive
Notes: … Firstly, [click] electrocardiogram monitoring (or ECG for short). …
﻿The ECG is a measure of the electrical current generated by the heart each heartbeat [click]. It is acquired by measuring the voltage difference between two points on the body surface over time [click]. Traditionally, it has been measured on the chest, close to the heart. Recently, chest-patch devices have been developed which measure the ECG between two electrodes in close proximity. In addition, wrist-worn devices [click] have been developed which can measure the ECG between opposite arms. The ECG can be used to diagnose a range of heart conditions.
Secondly, photoplethysmography [click] (or PPG for short). The PPG is a measure of changes in blood volume over time in a bed of tissue. It is usually acquired by illuminating the skin with an LED, and measuring the intensity of light either transmitted through or reflected from the skin. The PPG is widely measured by pulse oximeters at the finger, but many wearables [click] now measure it at the wrist. The PPG is used in wearables primarily for heart rate monitoring [click], but can also be used to assess oxygen saturation and several other parameters.
Thirdly, accelerometers [click] are used to detect and quantify movements. Accelerometers measure acceleration and are widely incorporated into wearables as they are cheap and consumer little power. Accelerometry can be used to count steps [click], identify falls, and as a reference motion signal with which to either identify or remove the effects of motion from other signals.
The functionality of wearables is mainly determined by their sensor technologies, and the algorithms used to derive parameters from the sensor signals.
… This table [click] summarises the functionality of a few selected consumer wearables. …

## Slide 8
The Photoplethysmogram
Photoplethysmogram (PPG) Sensor
Wrist cross-section
Notes: If you have a fitness tracker or smart watch you may have seen sensors on its underside, perhaps emitting red or green light. These are PPG sensors. I’ll use a cross-section [click] of the wrist to illustrate how a PPG signal is generated and captured. In this example a [click] green LED is used to illuminate the wrist. Some of the light is attenuated in the tissues, and some is reflected back towards the sensor. The level of reflected light is recorded by a [click] light sensor (in yellow). Here is [click] an example of the resulting PPG signal.
[click] The cross-section shows the pulsatile flow of arterial blood into the wrist with each heart beat. On the right, the signal increases and decreases with each heart beat, as the volume of arterial blood, which absorbs light, increases and decreases. This forms pulse waves – one for each heart beat.

## Slide 9
Photoplethysmogram (heart rate = 100 bpm)
Slower heart rate (heart rate = 80 bpm)
Atrial fibrillation (heart rate = 140 bpm)
Photoplethysmogram (PPG) Sensor
Wrist cross-section
The Photoplethysmogram
Notes: [click] Heart rate: 100 bpm
[click] Heart rate: 60 bpm
[click] irregular rhythm
[click] irregular intervals
[click] varying pulse amplitudes
If you have a fitness tracker or smart watch you may have seen sensors on its underside, perhaps emitting red or green light. These are PPG sensors. I’ll use a cross-section [click] of the wrist to illustrate how a PPG signal is generated and captured. In this example a [click] green LED is used to illuminate the wrist. Some of the light is attenuated in the tissues, and some is reflected back towards the sensor. The level of reflected light is recorded by a [click] light sensor (in yellow). Here is [click] an example of the resulting PPG signal.
[click] The cross-section shows the pulsatile flow of arterial blood into the wrist with each heart beat. On the right, the signal increases and decreases with each heart beat, as the volume of arterial blood, which absorbs light, increases and decreases. This forms pulse waves – one for each heart beat.
[click] This PPG signal shows a heart rate of approximately 100 beats per minute. [click] Here is an example of a slower heart rate, still in sinus rhythm. Besides the heart rate, the signal is also greatly influenced by the [click] heart rhythm, since an irregular rhythm such as atrial fibrillation results in both [click] irregular beat-to-beat intervals and [click] varying pulse amplitudes. [click] Focusing in on a [click] single pulse wave [click]...

## Slide 10
Photo by Luke Chesser on Unsplash
Using photoplethysmography to prompt single-lead ECG

## Slide 11
Single-lead ECG
“heart rhythm with no discernible repeating P waves and irregular RR intervals”
2020 ESC Guidelines for the diagnosis and management of atrial fibrillation…

## Slide 12
Source: SAFER Initial library of ECGs
Beat detection
Rhythm analysis
P-wave analysis
Rhythm analysis
Detecting AF from beat-to-beat intervals
Notes: In this talk I’ll provide an overview of the most promising opportunities to use photoplethysmography in public health. I’ll introduce photoplethysmography, [click] covering the physiological and technical fundamentals. I’ll then discuss [click] potential clinical applications, drawing on real-life examples of how they could be used to improve health outcomes. Finally, [click] I’ll highlight pressing directions for future research to ensure photoplethysmography can be used to inform clinical decisions safely and robustly.

## Slide 13
Measurement process
User wears device
Device measures signal(s)
Signal processing
Interpretation of measurements

## Slide 14
1. Introduction to wearables
2. Our contributions
Photo by Luke Chesser on Unsplash
3. Perspectives
Photo by Dario Valenzuela on Unsplash

## Slide 15
Detecting AF in daily life
User wears device
Device measures signal(s)
Signal processing
Interpretation of measurements
Older adults (aged 70+)
Continuous wrist photoplethysmography monitoring
Intermittent ECG measurements

## Slide 16
Aim: Assess performance and acceptability of wearables for detecting atrial fibrillation (AF)
Methods:
In 130 older adults, aged 65+, half of whom have AF
Two wrist-worn devices: smartwatch, wristband
Reference ECG chest patch
Questionnaire
Progress:
21 out of 130 participants to date.
The SAFER Wearables Study, NCT04715555

## Slide 17
The SAFER Wearables Study, NCT04715555

## Slide 18
The SAFER Wearables Study, NCT04715555
“looked as though I'd been attacked with an octopus with a round ring”
"I forget that's there”

## Slide 19
Detecting AF in daily life
User wears device
Device measures signal(s)
Signal processing
Interpretation of measurements
Beat detection

## Slide 20
Aim: To identify the best algorithm to detect heart beats in photoplethysmography signals.
Methods:
Fifteen open-source beat detection algorithms
Assessed against electrocardiogram-derived heartbeats
On eight datasets.
Charlton PH et al., Detecting beats in the photoplethysmogram: benchmarking open-source algorithms, Phys Meas., 2022. https://doi.org/10.1088/1361-6579/ac826d (CC BY 4.0)
Benchmarking photoplethysmogram beat detectors

## Slide 21
Charlton PH et al., Detecting beats in the photoplethysmogram: benchmarking open-source algorithms, Phys Meas., 2022. https://doi.org/10.1088/1361-6579/ac826d (CC BY 4.0)
Identified two best-performing algorithms to detect heart beats
Benchmarking photoplethysmogram beat detectors

## Slide 22
Benchmarking photoplethysmogram beat detectors
Charlton PH et al., Detecting beats in the photoplethysmogram: benchmarking open-source algorithms, Phys Meas., 2022. https://doi.org/10.1088/1361-6579/ac826d (CC BY 4.0)
Eight beat detectors performed well at rest in the absence of movement
F1 scores of >=90% on hospital data and wearable data

## Slide 23
Benchmarking photoplethysmogram beat detectors
Charlton PH et al., Detecting beats in the photoplethysmogram: benchmarking open-source algorithms, Phys Meas., 2022. https://doi.org/10.1088/1361-6579/ac826d (CC BY 4.0)
Eight beat detectors performed well at rest in the absence of movement
F1 scores of >=90% on hospital data and wearable data
Their performance was poorer during exercise:
F1 scores of 55%–91%

## Slide 24
Benchmarking photoplethysmogram beat detectors
Charlton PH et al., Detecting beats in the photoplethysmogram: benchmarking open-source algorithms, Phys Meas., 2022. https://doi.org/10.1088/1361-6579/ac826d (CC BY 4.0)
Eight beat detectors performed well at rest in the absence of movement
F1 scores of >=90% on hospital data and wearable data
Their performance was poorer during exercise:
F1 scores of 55%–91%
Performance was:
Poorer in AF
Poorer in neonates than adults
Not associated with ethnicity (Black compared with White)
Concluded that ‘MSPTD’ and ‘qppg’ performed best, although this is somewhat subjective.

## Slide 25
Aim: To identify the best-performing open-source QRS detector for use with telehealth ECGs
Methods:
18 open-source beat detection algorithms
Assessed against manual annotations
On six datasets.
Performance expressed as F1-score (ranges from 0 to 1, with 1 being best).
Benchmarking ECG beat detectors
Kristof F et al., QRS detection in single-lead, telehealth electrocardiogram signals: Benchmarking open-source algorithms, PLOS Digital Health, 2024. https://doi.org/10.1371/journal.pdig.0000538 (CC BY 4.0)

## Slide 26
Results:
Using ‘F1 score ≥0.96’ as the criteria for determining whether an algorithm performed well:
12 algorithms performed well on ECGs collected under clinical supervision.
Benchmarking ECG beat detectors
Kristof F et al., QRS detection in single-lead, telehealth electrocardiogram signals: Benchmarking open-source algorithms, PLOS Digital Health, 2024. https://doi.org/10.1371/journal.pdig.0000538 (CC BY 4.0)
Fewer performed well on telehealth ECGs:
  Five performed well on the TELE dataset
  Performance was poorer on low-quality SAFER data.
  Six performed well on high-quality SAFER data

## Slide 27
Detecting AF in daily life
User wears device
Device measures signal(s)
Signal processing
Interpretation of measurements
Automated AF detection
Manual ECG interpretation

## Slide 28
Detecting AF from beat-to-beat intervals
Aim:
To characterise RR intervals in AF detected through screening
Methods:
30-second ECGs collected from adults aged 65 and over in an AF screening study (the SAFER Feasibility Study)
2,709 ECGs were analysed:
  671 ECGs exhibiting AF (from 43 subjects)
  2,038 non-AF ECGs (each from a different subject)
Akande R et al., Characterising RR Intervals in Atrial Fibrillation Detected through Screening, CinC, 2023. https://doi.org/10.22489/CinC.2023.270 (CC BY 4.0)
Kristof F et al., QRS detection in single-lead, telehealth electrocardiogram signals: Benchmarking open-source algorithms, PLOS Digital Health, 2024. https://doi.org/10.1371/journal.pdig.0000538 (CC BY 4.0)

## Slide 29
Characterising RR-intervals in atrial fibrillation
All differed significantly between AF and non-AF
RRstd: standard deviation
RRvar: variability
RMSSD: root mean square of successive RR interval differences
SD1: standard deviation derived from a Poincare plot
pNN50: percentage of successive RR intervals that differ by more than 50ms
RRpm60: percentage of RR intervals within ±60ms of the median RR interval
AUROC: 96.3%
Akande R et al., Characterising RR Intervals in Atrial Fibrillation Detected through Screening, CinC, 2023. https://doi.org/10.22489/CinC.2023.270 (CC BY 4.0)

## Slide 30
Manual diagnosis of AF from single-lead ECGs
Hibbitt K et al., Reliability of single-lead electrocardiogram interpretation to detect atrial fibrillation: insights from the SAFER feasibility study, EP Europace, 2024. https://doi.org/10.1093/europace/euae181 (CC BY 4.0)
In the SAFER feasibility study, 1,843 single-lead ECGs from 185 participants were independently reviewed by two cardiologists
Moderate agreement between cardiologists on participant-level diagnoses.
For every 100 agreed diagnoses, there would be 70 disagreements.
Chart: 
| 1st Qtr | 2nd Qtr
Sales | 100 | 70
44 participants with
agreed AF diagnosis
31 participants with complete disagreement over AF diagnosis
The quality of ECG signals greatly influenced the reliability of single-lead ECG interpretation.

## Slide 31
1. Introduction to wearables
2. Our contributions
Photo by Luke Chesser on Unsplash
3. Perspectives
Photo by Dario Valenzuela on Unsplash

## Slide 32
Perspectives
Open questions:
The roles of traditional signal processing (e.g. using hand-crafted features) vs. deep learning.
How important is the accuracy of wearable data and subsequent diagnsoses?
What level of AF burden warrants treatment?

## Slide 33
Prof Jonathan Mant
Prof Panicos Kyriacou
Rayo Akande
Joachim Behar
Timothy Bonnici
James Brimicombe
Cheryl Chapman
Kelly Ding
Andrew Dymond
Miranda Van Emmenis
Katie Hibbitt
Florian Kristof
Vaidotas Marozas
Elisa Mejia-Mejia
Kate Williams
The SAFER Research Team
University of Cambridge
City, University of London
King’s College London
British Heart Foundation
NIHR
EPSRC
With thanks to…
… and many others

## Slide 34
Sonia Sevilla, https://commons.wikimedia.org/wiki/File:Kamar_Zard_Buzhan_-_Nishapur_1.jpg (CC0 1.0)
Potentially, wearables provide a holistic approach to detect atrial fibrillation, from unobtrusive rhythm monitoring to acquisition of diagnostic ECGs.
There is much work to be done to optimise device acceptability and performance, and make wearables as reliable as a climbing rope.

## Slide 35: Towards accurate wearable data for clinical decision making
Photo by Luke Chesser on Unsplash
Slides available at: https://doi.org/10.5281/zenodo.13347570 (CC BY 4.0)
Dr Peter H. Charlton
https://peterhcharlton.github.io
pc657@cam.ac.uk

## Slide 36
Charlton PH et al., Wearable photoplethysmography for cardiovascular monitoring, Proc. IEEE, 2022. ﻿https://doi.org/10.1109/JPROC.2022.3149785 (CC BY 4.0)
Adapted from: ﻿H. Gray, Anatomy of the Human Body, W. H. Lewis, Ed., 2nd ed. New York, NY, USA: Lea and Febiger, 1918
Photoplethysmogram (PPG) Sensor
Wrist photoplethysmography: limitations

## Slide 37
Peter Charlton, Wikimedia Commons, https://commons.wikimedia.org/wiki/File:12_lead_ECG.jpg (CC BY 4.0)
“Sinus P waves are usually most prominently seen in leads II and V1”
Meek S and Morris F, ‘ABC of clinical electrocardiography: Introduction. II—basic terminology’, https://doi.org/10.1136/bmj.324.7335.470
Dry electrodes
Recorded without clinical supervision
Lead I
Single-lead ECG: limitations

## Slide 38
Older adults because:
AF incidence increases with age
Wu J et al., ‘Temporal trends and patterns in atrial fibrillation incidence: A population-based study of 3.4 million Individuals’, https://doi.org/10.1016/j.lanepe.2022.100395 (CC BY 4.0)
In daily life because:
AF can occur infrequently
From primary and secondary electronic health records of 3.4 million individuals in England
Target population and target setting
Notes: NICE (2021) recommends offering anticoagulation “to people with atrial fibrillation and a CHA2DS2‑VASc score of 2 or above” (i.e. after detecting AF, all women ≥ 65 years, and all men ≥ 75 years).

## Slide 39
Not everyone has a smartwatch, particularly older adults
Shandhi MH et al., ‘Assessment of ownership of smart devices and the acceptability of digital health data sharing’, https://doi.org/10.1038/s41746-024-01030-x
(CC BY 4.0)
18-25
26-41
42-57
58-76
77+
Age
0
25
50
75
100
Ownership Percentage (%)
An electronic survey of 1,368 patient advisory group members:
Smartwatch ownership