## Slide 1: An assessment of algorithms to estimate respiratory rate from the electrocardiogram and photoplethysmogram
P. H. Charlton and T. Bonnici, L. Tarassenko, D. A. Clifton, R. Beale and P. J. Watkinson
DOI: 10.1088/0967-3334/37/4/610

## Slide 2: Respiratory Rate
The most sensitive marker of clinical deterioration
Notoriously poorly recorded
  Missing
  Inaccurate
Difficult to measure manually
Thoracic bands uncomfortable

## Slide 3: Literature
Over 100 RR algorithms
Not possible to compare algorithms using the published results
Limitations:
  No standard algorithm implementations for benchmarking
  Atypical populations - ventilated subjects, children
  Different statistical measures
  No compensation for repeated measures

## Slide 4: Aims
Identify which algorithm performs the best using appropriate statistical measures
Contextualise algorithm performance by comparing with the current non-invasive standard, impedance pneumography
Compare performance when using ECG or PPG
Provide a benchmark toolbox of algorithms and data for the benefit of other researchers

## Slide 5: Prior Work
Notes: Image from: https://openclipart.org/detail/66667/publication [500 px PNG]

## Slide 6: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR

## Slide 7: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
PPG
No mod
BW
AM
FM
ECG

## Slide 8: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
PPG
No mod
BW
AM
FM

## Slide 9: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
PPG
No mod
BW
AM
FM
Identify fiducial points

## Slide 10: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
PPG
No mod
BW
AM
FM
Find baseline

## Slide 11: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
PPG
No mod
BW
AM
FM
Find baseline

## Slide 12: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
PPG
No mod
BW
AM
FM
Measure amplitudes and intervals

## Slide 13: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
PPG
No mod
BW
AM
FM
Obtain respiratory signals

## Slide 14: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
PPG
No mod
BW
AM
FM
breaths
14 techniques implemented

## Slide 15: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
12 techniques implemented
Notes: Ref = 20 bpm

## Slide 16: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
4 techniques implemented

## Slide 17: Structure of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
1 technique implemented
Fusion

## Slide 18: Constructing Algorithms
Fourier Transform
Autoregression
Peak detection
Zero-crossings
…
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
BW
AM
FM
Peak amplitudes
Onset amplitudes
…

## Slide 19: Constructing Algorithms
Fourier Transform
Autoregression
Peak detection
Zero-crossings
…
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
BW
AM
FM
Peak amplitudes
Onset amplitudes
…

## Slide 20: Constructing Algorithms
Fourier Transform
Autoregression
Peak detection
Zero-crossings
…
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
BW
AM
FM
Peak amplitudes
Onset amplitudes
…

## Slide 21: Constructing Algorithms
BW
AM
FM
Peak amplitudes
Onset amplitudes
…
Fourier Transform
Autoregression
Peak detection
Zero-crossings
…
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR

## Slide 22: Constructing Algorithms
Fourier Transform
Autoregression
Peak detection
Zero-crossings
…
BW
AM
FM
Peak amplitudes
Onset amplitudes
…
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR

## Slide 23: Constructing Algorithms
Fourier Transform
Autoregression
Peak detection
Zero-crossings
…
BW
AM
FM
Peak amplitudes
Onset amplitudes
…
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
Smart Fusion
Temporal Fusion
…
370 algorithms implemented

## Slide 24: Toolbox of Algorithms
Extraction of Respiratory Signals
RR Estimation
Fusion of RR Estimates
ECG or PPG
RR
Publicly available here
Charlton P.H. et al. Waveform analysis to estimate respiratory rate, in Secondary Analysis of Electronic Health Records, Springer, pp.377-390, 2016.
DOI: 10.1007/978-3-319-43742-2_26 . CC BY-NC 4.0 Licence

## Slide 25: Methods

## Slide 26: Verification of Implementations
Synthetic ECG and PPG with simulated RR modulation
  HR: 30-200 bpm
  RR: 4-60 bpm
314 (85%) of algorithms accurate
Failures caused by two techniques which were removed
Notes: We created a library of synthetic ECG and PPG modulated by RR such that we had samples with HRs ranging from 30-200bpm and RR ranging from 4-60bpm.

## Slide 27: Participants
Aged 18 to 40
Free of comorbidities affecting cardiac, respiratory or autonomic nervous systems
Range of RR generated by asking subjects to exercise
National Clinical Trial 01472133
Walk
2 min
Run
~ 5 min
Rest
10 min
Recover
10 min
Notes: Images from:
https://openclipart.org/detail/28443/sleeping-symbol
https://openclipart.org/detail/201733/walking-icon
https://openclipart.org/detail/223753/woman-running
All 250 px PNG

## Slide 28: Signals
Notes: 500Hz
RR from IP provided by monitors at 1Hz

## Slide 29: Signal Quality
high
high
low
low
ECG
PPG
Time
Time
Template
Beats
Notes: 32s windows
For each high quality window a template beat is constructed.
The correlation between individual beats and the template is assessed. If the correlation is below an empirically determined threshold, then the window is deemed to be of low quality

## Slide 30: Reference RRs
Positive-gradient crossings detected from oro-nasal pressure signal
Algorithm verified by comparison with manually annotated breaths

## Slide 31: Statistics
Consistent interpretation in different populations
Intuitive interpretation conducive to decision making
Separates bias from precision
  Trends are more important than absolute values
  If error is caused by a constant bias can be corrected by calibration

## Slide 32: Statistics
2SD
Correction for repeated measures using a random effects model
Coverage Probability
Ranked algorithms by 2SD, followed by bias.
Bias
Notes: The differences between the reference and estimated RR are plotted against the averages of the two techniques.
The bias is the mean difference.
The limits of agreement are the bias +/- 2 standard deviations. 95% of measurements fall within the LoA. The greater the value of 2SD, the wider the LoA, the less precise the measurements.
When calculating the LoA we corrected for repeated measures using a random effects model.
We also reported a complementary statistic, the coverage probability. This tells you what proportion of measurements fall within an acceptable error range. We set acceptable error at +/- 2bpm.

## Slide 33: Results

## Slide 34: Dataset
≈ 36 windows per subject
39 subjects
  Age: 29 (26, 32)
  BMI: 23 (21, 26)
  54% female

## Slide 35: Dataset
≈ 36 windows per subject
Respiratory Rate [bpm]
Heart
Rate
[bpm]
5
32
111
41
39 subjects
  Age: 29 (26, 32)
  BMI: 23 (21, 26)
  54% female

## Slide 36: Dataset
≈ 36 windows per subject
Respiratory Rate [bpm]
Heart
Rate
[bpm]
5
32
111
41
39 subjects
  Age: 29 (26, 32)
  BMI: 23 (21, 26)
  54% female
Publicly available here

## Slide 37: Performance of Algorithms

## Slide 38: Performance of Algorithms
Three techniques

## Slide 39: Performance of Algorithms

## Slide 40: Performance of Algorithms
Time
Freq

## Slide 41: Best Algorithms
Signal | Rank | 2SD [bpm] | RR Estimation | Modulation Fusion? | Temporal Fusion?
Clinical (IP) | 5 | 5.4 |  |  | 
ECG | 1 | 4.7 | Time | ✓ | 
 | 2 | 5.2 | Time | ✓ | 
 | 3 | 5.2 | Time | ✓ | 
 | 4 | 5.3 | Time | ✓ | 
 | 6 | 5.6 | Time |  | 
PPG | 15 | 6.2 | Time | ✓ | 
 | 17 | 6.5 | Time | ✓ | 
 | 35 | 7.0 | Time | ✓ | ✓
 | 46 | 7.5 | Time |  | ✓
 | 48 | 7.6 | Time |  | ✓
Same Algorithm

## Slide 42: ECG vs PPG
Significant difference in 2SD (median):
  ECG: 11.6 bpm
  PPG: 12.4 bpm
64% of algorithms more precise on ECG
Different physiological mechanisms

## Slide 43: Discussion

## Slide 44: Limitations
Not all algorithms implemented
Invite contributions
Statistics based on normally distributed errors
Cannot extrapolate to other scenarios

## Slide 45: Future Work
Investigate effects of:
Physiology
Signal Acquisition Equipment
RR Algorithm
Patient
RR
Age
Disease
Arrhythmias
Signal Fidelity
Filtering
Noise
This paper

## Slide 46: Future Work
Investigate effects of:
Physiology
Signal Acquisition Equipment
RR Algorithm
Patient
RR
Age
Disease
Arrhythmias
Signal Fidelity
Filtering
Noise
Charlton P.H. et al. Extraction of respiratory signals from the electrocardiogram and photoplethysmogram: technical and physiological determinants, Physiological Measurement, 37(4), 2016. DOI: 10.1088/1361-6579/aa670e . CC BY 3.0 Licence
This paper

## Slide 47: Future Work
Investigate effects of:
Physiology
Signal Acquisition Equipment
RR Algorithm
Patient
RR
Age
Disease
Arrhythmias
Signal Fidelity
Filtering
Noise
Charlton P.H. et al. Extraction of respiratory signals from the electrocardiogram and photoplethysmogram: technical and physiological determinants, Physiological Measurement, 37(4), 2016. DOI: 10.1088/1361-6579/aa670e . CC BY 3.0 Licence
This paper

## Slide 48: Conclusions
314 algorithms assessed under ideal conditions
According to these results …
  time-domain RR estimation, and
  fusion of estimates
  … resulted in superior performance.
Four ECG-based algorithms comparable to clinical standard
ECG preferable to PPG
Toolbox of algorithms and dataset publicly available

## Slide 49: Acknowledgments
The authors are grateful to …
Data collection:
J Brooks, I Schelcher, R Yang, K Lei and J Smith
Algorithm implementations:
M Pimentel and C Orphanidou
Statistical analysis:
J Birks and S Gerry
Funders:
EPSRC, NIHR, Wellcome Trust, Royal Academy of Engineering
The views expressed are those of the authors and not necessarily those of the EPSRC, NHS, NIHR, Department of Health, Wellcome Trust, or Royal Academy of Engineering.
A complete list of acknowledgments is available here.

## Slide 50: Additional Acknowledgments
Thanks also to:
Jason Long for Cayman Theme which inspired this presentation template
Open Clipart for some of the images in this presentation
Source:
This presentation was adapted from previous presentations by P. H. Charlton which are publicly available under the Creative Commons Attribution 4.0 Licence. DOIs: 10.5281/zenodo.166525 and 10.5281/zenodo.166546 .

## Slide 51: References and Resources
Part of the Respiratory Rate Estimation Project at:
http://peterhcharlton.github.io/RRest/
The dataset is available here.
The algorithms and user manual are available here.
The complete table of results is in the Supplementary Material
A complete list of references is available here.
Charlton P.H. and Bonnici T. et al. An assessment of algorithms to estimate respiratory rate from the electrocardiogram and photoplethysmogram,
Physiological Measurement, 37(4), 2016.
DOI: 10.1088/0967-3334/37/4/610 . CC BY 3.0 Licence